Performance Maintenance Protocol

2690/2695/e2695 Separations Module

| Waters system: |  |
| --- | --- |
| PM date: |  |

# Guidelines

This performance maintenance (PM) protocol outlines maintenance tasks you should regularly perform on Waters systems. Regular maintenance ensures optimum performance and uninterrupted trouble-free operation.

## Schedule

Waters recommends that performance maintenance visits be performed once a year. Systems under heavy use or using aggressive solvents need more frequent maintenance. For detailed maintenance schedules, refer to the appropriate user guide located on Waters.com. Failure to follow the maintenance schedules as described in the user guide might result in unreliable operation and inaccurate laboratory results.

## Safety considerations

Some protocols might involve removing instrument covers, exposing the performer to high voltages, high currents, and moving parts. This presents risks of electric shock or damage to the instrument. To reduce these risks, Waters recommends that these protocols be conducted by certified Waters field service engineers with access to the latest information on Waters products. Waters is not responsible for actions initiated by an unqualified performer. If you are unsure, have a Waters certified field service engineer conduct the performance maintenance activity.

## Waters Quality Parts

Waters designs and manufactures Waters Quality Parts® to the same strict regulatory standards it uses for its high-performance Waters systems. Waters recommends using Waters Quality Parts for all maintenance activities. Failure to use Waters Quality Parts might compromise laboratory results and reliability.

## Customer education

Waters offers courses on operating and maintaining its high-performance systems, as well as courses designed to help you maintain professional competency and further your knowledge in chromatography, mass spectrometry, and Waters data systems. For more information on available training courses, visit Waters Educational Services on Waters.com.

## Qualification/compliance services

Waters Compliance Services help you achieve and maintain the regulatory compliance of your LC and MS instruments and software with a combination of expertise, compliant-ready systems and software, and a portfolio of services to ensure that you consistently obtain quality data. To learn more, visit Qualification/Compliance Services on Waters.com.

# PM task overview

Table 1 is an overview of the PM tasks performed.

Note:  For instructions on performing these procedures, refer to the Waters Alliance e2695 Separations Module Operator’s Guide.

Table 1 - PM tasks

| Inspect | Replace/rebuild | Verify | Optional calibration tests |
| --- | --- | --- | --- |
| Firmware version | Filters | Initialization diagnostics | Flow rate accuracy |
| Lifetime counter | Solvent management components | Pressure test | Gradient proportioning valve |
| Solvent manager door filter | Sample management components | Adjust the seals | Injection accuracy |
| Fan | Lithium battery | Compression check |  |
| Valves |  |  |  |
| Carousel drive mechanism |  |  |  |
| PerformancePlus™ degasser |  |  |  |
| Seal wash system |  |  |  |
| Fittings and tubing (leaks) |  |  |  |
| Error logs |  |  |  |

# PM task descriptions

Performance maintenance consists of three task categories:

Inspect

Replace or rebuild parts

Verify

## Note:  If the instrument fails pre- or post-performance maintenance inspections or verifications, it might need additional service beyond the scope of this protocol. Such service might incur additional parts and labor charges.

## Inspect

- If the instrument is off, power it on and allow it to complete software checks.

- Inspect the firmware and, if appropriate, update it.

- Inspect the lifetime counter and record the eluent delivered and injection count.

- Ensure that the solvent manager door filter is clear of debris and dust; clean if necessary.

Ensure that the fan grate on the rear of the instrument is clear of debris and dust; clean if necessary.

Inspect the following for leaks, precipitation, corrosion, proper operation, or any obvious damage:

Valves – prime/vent; high pressure 1 and 2; low pressure 3; gradient proportioning

Carousel drive mechanism

PerformancePlus degasser

Seal wash system

Fittings and tubings

Review error logs for any errors and troubleshoot as necessary.

## Replace or rebuild parts

Using the parts in the PM kit, replace or rebuild parts as necessary.

Note:	For instructions on replacing parts, see the appropriate user manual on Waters.com.

## Verify

After completing the maintenance tasks, verify instrument performance:

- Power-on the instrument and allow it to complete on-board software checks.

- Using Methanol at 1mL/min and a back pressure of at least 1700 psi, inspect the delta pressure and record results. If delta is larger than 2.5% of the system pressure, run the static leak test to locate a possible leak.

Adjust the seals from the Diagnostics menu and record the results.

Perform a compression check and record the results.

Note:	For information on priming all solvent lines and purging injectors, refer to the Waters Alliance e2695 Separations Module Operator’s Guide.

# Results/checklist

A Waters field service engineer completes the test results, checklist, and confirmation sections.

## Test results

Table 2 - Firmware version

| Firmware revision |  |
| --- | --- |

Table 3 – Lifetime counter results

|  | Lifetime counters |
| --- | --- |
| Eluent delivered |  |
| Injection count |  |

Table 4 – Pressure test

| System pressure | Delta pressure | Acceptable delta range | Pass or fail |
| --- | --- | --- | --- |
|  |  | < 2.5% system pressure | Pass Fail |

Table 5 - Adjust seals test results

| Seal data | New | Previous | Acceptable range | Pass or fail |
| --- | --- | --- | --- | --- |
| Top |  |  | 30 to 80 | Pass Fail |
| Bottom |  |  | 80 to 200 | Pass Fail |
| Width |  |  | 50 to 120 | Pass Fail |

Table 6 - Compression check test results3

| Parameter | Value | Pass or fail |
| --- | --- | --- |
| Compression volume | µL | Pass Fail |
| Delta decay pressure | psi | Pass Fail |
| Final pressure | psi | Pass Fail |

Table 7 - Optional calibration test results – flow rate accuracy

|  | Timer |
| --- | --- |
| Serial/ID number |  |
| Calibration date |  |

Table 8 - Optional calibration test results – flow rate accuracy

| Value | Pass or fail |
| --- | --- |
| min Expected result: 2:51 to 3:09 minutes | Pass Fail |

Table 9 - Optional calibration test results – gradient proportioning valve

| Solvent | Measured volume | Pass or fail |
| --- | --- | --- |
| A | mL | Pass Fail |
| B | mL | Pass Fail |
|  | Expected result: B cylinder volume = A cylinder volume ±3 mL |

Table 10 - Optional calibration test results – injection accuracy

|  | Balance |
| --- | --- |
| Serial/ID number |  |
| Calibration date |  |

Table 11 - Optional calibration test results – injection accuracy

| W1 | W2 | Programmed volume (µL) | Calculated volume (µL) | Pass or fail |
| --- | --- | --- | --- | --- |
| g | g | 50 µL | µL Expected result: 50 µL± 2.0 µL | Pass Fail |
| g | g | 12 µL | µL Expected result: 12 µL± 1.0 µL | Pass Fail |

## Checklist

Note: Select N/A if the task does not apply.

### Inspect

| All inspect tasks completed | Yes No N/A |
| --- | --- |
|  | Comments: |

### Replace/Rebuild

| Filters | Yes No N/A |
| --- | --- |
| Solvent management components | Yes No N/A |
| Sample management components | Yes No N/A |
| Lithium battery | Yes No N/A |
|  | Comments: |

### Verify

| All verify tasks completed | Yes No N/A |
| --- | --- |
|  | Comments: |

# Performance maintenance confirmation

| I confirm on behalf of |  |  |  |
| --- | --- | --- | --- |
|  | Company name |
| that this performance maintenance procedure has been satisfactorily completed on: |  |
|  |  | Serial no. |
| Signed: |  | Date: |
| Customer |  |
| Print name |  |
| I certify that the maintenance procedure has been successfully completed. |  |
| Signed: |  | Date: |
| Waters engineer |  |
| Print name |  |
| Comments: |

| Waters Corporation 34 Maple Street Milford, Massachusetts 01757 USA |  |
| --- | --- |
| PerformancePlus , The Science of What’s Possible, Waters, and Waters Quality Parts are trademarks of Waters Corporation. All other trademarks are property of their respective owners. ©2020 Waters Corporation. Printed in the U.S.A. |

# Footnotes

Waters field service engineers perform optional calibration tests as part of a Total Assurance with Calibration Plan. For an additional charge, Waters can   perform these tests for customers without this plan. Some of these tests require chemistry kits not supplied in the performance maintenance kit. For details on performing the optional calibration tests, see the appropriate user guide on Waters.com.

Only Waters field service engineers can replace batteries.

For information on adjusting the seals and checking compression, refer to the Waters Alliance e2695 Separations Module Operator’s Guide.